## Slide 1: Twin Probe design
M W Hutchings August 2019

## Slide 2: twin probe design
Design
Outer Tube
Inner Rod
Insulation
Fluid
Probe P1
Probe P2
P1
P2
Air

## Slide 3: Basic Theory behind algorithms
Capacitance calculation for simple coaxial cable
Where Ɛ is permittivity, l is length, R2 and R1 are radius as shown below

## Slide 4: Basic Theory behind algorithms
Insulator has effect of giving two different permittivities
  2.05 for PTFE Probe Insulation
  2.25 for diesel fuel
  Air gap between conductor and PTFE Probe Insulation can also be considered
  This complicates previous equation and gives an expression of the form

## Slide 5: Basic Theory behind algorithms
Abbreviations:
  P0E Electronics reading at calibration
  P1E Probe 1 empty reading at calibration
  P1F Probe 1 full reading at calibration
  P2E Probe 2 empty reading at calibration
  P2F Probe 2 full reading at calibration
  P0 Live electronics reading
  P1 Live probe 1 reading
  P2 Live probe 2 reading

## Slide 6: Basic Theory behind algorithms
Single probe (no temperature compensation or dielectric compensation)
Single probe including temperature compensation (no dielectric compensation)

## Slide 7: Basic Theory behind algorithms
Existing twin probe design including temperature and dielectric compensation:
  Note: Highlighted figures were all empirically derived from test data.
  1.40 / 1.05 constants relate probe changes with temperature to electronics changes with temperature
  14.4 / 7.07 and 6.3 constants used multi fluid empirical data from code validation testing. Excel file is ‘results 8r funct testing cal in diesel’

## Slide 8: Basic Theory behind algorithms
Simplified in more meaningful terminology:

## Slide 9: Basic OPERATION
Basic Operation:
  Three capacitive readings taken. Two probes and the electronics
  Probe readings corrected for temperature by monitoring electronics addition
  Two probes ratio over full span at calibration in diesel used to calculate ‘m’
  ‘m’ used to calculate ‘divisor’ at calibration.
  Live divisor calculated in new fluid
  Divisor indicates Dielectric Constant of fluid
  Divisor and capacitance output from P1 probe used to calculate fluid level
  Divisor is calculated continually but some filtering to prevent sloshing effect
  Every 15 minutes divisor saved to EPROM
  Below 20% level the divisor is frozen. Stops large errors due to small signals